## Announcement:
differential pressure transmitter PS 27

Compact differential pressure transmitter for basic applications

- Measurement ranges from 0..50 Pa to 100 kPa

- Either with one fixed measurement range or toggling between four different measurement ranges

- For measuring positive or negative pressures, even asymmetrically

- Also available with two-wire technology (optional)

- Measurement ranges and output signals can be selected via jumpers (optional)

- Can be calibrated

- With 1 alarm contact (optional)

- With LCD (optional)

- Top-hat rail mounting and wall surface installation IP 65

Technical data

Measurement ranges:	50/100/200/500 Pa, 1/2/5/10 kPa

Analog outputs: 		0 .. 10 V, 2 .. 10 V, 0 .. 20 mA, 4 .. 20 mA, 0 .. 5 V

Power supply:	24 V AC / DC ± 10 % (without galvanic separation),
15 .. 32 VDC two-wire

LCD: 		3 ½-digits (optional)

Contact points: 	1 relay (changeover contacts), max. 230 V AC, 6 A
(min. required switching capacity 300 mW) (not for two-wire)

Measurement accuracy: 	± 3% FS for < 100 Pa or 
± 2 % FS for ≥ 100 Pa

Temperature coefficient zero point:	max. ±0.1 % FS / K

Max. system pressure:	10 kPa

Overload capacity:	12 x for measurement ranges ≤ 20 kPa
4 x for measurement ranges ≥ 20 kPa

Medium: 		dry air, non-aggressive and non-flammable gases

Step response time (T63)

(Time constant):		20 ms .. 4 s adjustable (factory-provided)

Rated temperature range:	- 20 .. 60 ° C with display 0 .. 50 ° C

Storage temperature: 	- 20 .. 70 ° C

Power consumption: 	max. 1 VA

Weight:		approx. 250 g

Cable glands:		2 x M12 clamp range Ø 3 .. 6.5 mm

Pressure connections:	tube fitting 4 mm / 6 mm or
Plug-in fitting 4 mm or
Plug-in fitting 6 mm

Protection class:		IP 65

Housing/ cover:		ABS

Dimensions:		80 mm x 87 mm x 52 mm

Wall mounted housing: 	Top-hat rail assembly and wall mounting possible

Certificates:		CE / UKCA

Optional:		DAkkS calibration certificate, ISO factory calibration certificate

Designation:	differential pressure transmitter

Type: 	PS 27

Manufacturer: 	halstrup-walcher GmbH